This experiment was performed on Siemens ONCOR linear accelerator for both x-rays beam energy using 3 dimensional water phantom (Blue phantom, IBA Germany). The dimension of water tank is 480 mm x 480mm x 400 mm and walls are made of acrylic. The point accuracy of water phantom in 0.1 mm having 500 mm/s scanning speed. Standard relative dosimetry setup was arranged for measurement. For physical wedge factor (PWF), we use CC13 ion chambers, (IBA, Germany) along with portable IBA electrometer CU500E and for virtual wedge factors (VWF), LDA 99 with emxx electrometer. Both electrometers for wedge factors dosimetry were connected through computer having Omnipro-accept software to obtain observations. Ion chamber was kept at beam’s central axis, with chamber center at water surface, such that the distance from source to surface (SSD) was 100cm.The wedge angles for both physical wedge and virtual wedges on which the wedge factors were calculated are 15, 30, 45 and 60.The depth remains 10 cm during our experiment. This work was performed for three different field sizes (10 x 10 cm2, 15 x 15 cm2 and 20 x 20 cm2) and for 15 MV and 6 MV energies.